## Sheet1
Name | The trackable stylet fit as well within the introducer needle as the solid stylet | It was as easy to withdraw the trackable stylet from the introducer needle as the solid stylet | The fibre-optic cable between the trackable stylet and the tracking system did not affect my ability to carry out the procedure | The fibre-optic cable between the trackable stylet and tracking system was the right length | The presence of the tracking array on the imaging probe did not affect my ability to carry out the procedure | Please describe your level of experience with ultrasound-guided needle procedures, and biopsies specifically. You may include your grade, for example "ST1". | The ultrasound imaging feed was displayed on the tracking system with suitable image quality | Any noticeable latency on the imaging feed did not affect my ability to carry out the procedure | The tracking visualisation was clearly visible on top of the ultrasound imaging feed | The tracking cursor did not obstruct the ultrasound image in a way that hindered my ability to perform the procedure | I could easily interpret the distance of the needle tip from the imaging plane from the tracking visualisation | The presence of the tracking visualisation made it easier for me to carry out the procedure | Any noticeable latency of the displayed tracked location did not affect my ability to carry out the procedure | The update rate of the tracking cursor was appropriate for the procedure | The depth of the needle tip within the phantom did not negatively affect the tracking accuracy | The insertion angle of the needle did not negatively affect the tracking accuracy | The tracking was accurate enough to aid with performing the procedure
| Agree | Agree | Agree | Agree | Agree | Consultant interventional radiology | Agree | Agree | Agree | Strongly agree | Agree | Strongly agree | Agree | Agree | Agree | Agree | Agree
| Agree | Agree | Agree | Agree | Agree | Consultant Interventional Radiologist. High level of experience with ultrasound-guided needle procedures. | Agree | Agree | Agree | Agree | Agree | Agree | Agree | Somewhat agree | Agree | Agree | Agree
| Agree | Agree | Agree | Agree | Agree | ST2, biopsies done on phantoms only, few US guided drain insertions. | Agree | Agree | Agree | Somewhat agree | Neutral | Agree | Neutral | Somewhat agree | Agree | Somewhat agree | Agree
| Agree | Agree | Agree | Agree | Somewhat agree | ST1 I have done some ascitic and pleural drains under supervision but no biopsies. | Agree | Neutral | Somewhat agree | Agree | Somewhat agree | Agree | Disagree | Somewhat agree | Disagree | Somewhat agree | Somewhat agree
| Agree | Agree | Agree | Agree | Somewhat disagree | St5 ir - 200 us procedures | Agree | Agree | Agree | Agree | Agree | Agree | Agree | Neutral | Somewhat disagree | Agree | Agree
| Somewhat agree | Somewhat agree | Somewhat agree | Somewhat agree | Somewhat agree | Extensive experience (approx 400 image guided biopsies), consultant level | Agree | Agree | Agree | Somewhat agree | Somewhat agree | Neutral | Somewhat agree | Somewhat agree | Somewhat disagree | Neutral | Somewhat agree
| Agree | Agree | Agree | Agree | Agree | clinical Fellow/SpR | Agree | Agree | Agree | Agree | Somewhat agree | Neutral | Agree | Agree | Somewhat agree | Agree | Somewhat agree
| Agree | Agree | Agree | Agree | Somewhat agree | ST4, I've done about 20 US guided needle procedures | Agree | Agree | Agree | Agree | Somewhat agree | Agree | Somewhat agree | Agree | Disagree | Agree | Agree
| Somewhat agree | Neutral | Somewhat agree | Somewhat agree | Somewhat agree | I have been performing these independently for 5 years | Agree | Agree | Somewhat agree | Agree | Agree | Somewhat agree | Agree | Agree | Agree | Agree | Agree
| Agree | Somewhat agree | Agree | Somewhat agree | Somewhat agree | Post CCT | Agree | Somewhat agree | Somewhat agree | Neutral | Somewhat agree | Somewhat agree | Somewhat agree | Agree | Somewhat agree | Agree | Somewhat agree
| Agree | Agree | Agree | Agree | Agree | Consultant | Agree | Somewhat agree | Agree | Agree | Agree | Agree | Agree | Agree | Agree | Agree | Agree
| Agree | Agree | Agree | Agree | Somewhat disagree |  | Agree | Somewhat agree | Agree | Agree | Agree | Agree | Somewhat agree | Agree | Somewhat agree | Neutral | Somewhat agree
| Stylet Fit in Needle | Stylet Withdrawal Ease | Cable Annoyance | Cable Length | Array Annoyance |  | Image Quality | Satisfactory Latency | Track Clarity | Cursor Annoyance | Elevation Clarity | Tracking Benefit | Satisfactory Latency | Satisfactory Rate | Depth Effect | Angle Effect | Satisfactory Accuracy
Strongly agree | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0
Agree | 10 | 9 | 10 | 9 | 5 | 0 | 12 | 8 | 9 | 8 | 6 | 7 | 6 | 7 | 5 | 8 | 7
Somewhat agree | 2 | 2 | 2 | 3 | 5 | 0 | 0 | 3 | 3 | 2 | 5 | 2 | 4 | 4 | 3 | 2 | 5
Neutral | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 2 | 0
Somewhat disagree | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0
Disagree | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0
| 12 | 12 | 12 | 12 | 12 | 0 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12